Nucleic Acid Quality Control

Code: EGTP-CAN

Taper: Analysis

Prix: From 500 DZD

Description:

Nucleic Acid Quality Control

Code: EGTP-CAN | Type: Analysis | Price: 500 DZD per sample

Service overview

Assessment of purity, concentration, and integrity of nucleic acids (DNA/RNA) to ensure suitability for downstream applications (PCR, cloning, sequencing). Measurements are performed by spectrophotometry (ScanDrop) and/or agarose gel electrophoresis as requested.

Sample requirements (critical)

- DNA: ≥20 µL at 5–100 ng/µL in nuclease-free water or TE.

- RNA: ≥20 µL at 10–100 ng/µL in RNase-free water/TE; keep on ice/≤4 °C and ship promptly.

- Provide for each sample: internal code, nucleic acid type (DNA/RNA), source/organism, extraction method/kit, and any additives (e.g., EDTA, salts).

Workflow & deliverables

- Spectrophotometry (ScanDrop):

  - Concentration (A260), A260/280 and A260/230 purity ratios.

- Agarose gel electrophoresis (optional/if requested):

  - Integrity/fragmentation check and approximate size profile.

- Deliverables: 1-page QC report (per sample) with measured values, pass/fail comments for common downstream uses, and gel image when applicable.

Turnaround & pricing notes

- Turnaround: typically 2–5 working days from sample receipt.

- Pricing: 500 DZD per sample (spectrophotometry only).

  - Add +500 DZD if gel integrity check is requested (optional).

  - Large batches may be scheduled based on instrument availability.

Limitations

- Spectrophotometry can overestimate concentration in the presence of contaminants (phenol/salts); if high accuracy is required (e.g., NGS), request a fluorometric assay (quoted separately).

- Degraded or contaminated samples may be flagged as not recommended for downstream applications.

How to order

Submit the EGTP-CAN request indicating DNA or RNA, whether you require gel assessment, and the intended downstream application (PCR, Sanger, WGS) so we can apply appropriate acceptance thresholds.